## Calculator
Noise Figure Measurement Uncertainty Calculator
© Keysight Technologies, Inc. 1999, 2012, 2015-6
Differences from legacy calculator |  | Other NFUCs
|  |  |  | Specification Style | Distribution | Calculations Zone |  | Sensitivities
Device Under Test | DUT NF, NF1 | 3 | dB | Fixed |  | F1, Power Ratio | 1.9953 | F12/F1= | 1.0451 |  |  |  |  | No calibration case: combining
| DUT Gain, G1 | 20 | dB | Fixed |  | G1, Power Ratio | 100 | F2/F1G1= | 0.050119 |  | Sens, int cal | 0.045107 |  | 2nd stage bias with Gaussian-
| DUT Input Match* | 1.5 | VSWR | Fixed | Fixed | S11, ρ | 0.2 | (F2-1)/F1G1= | 0.045107 |  |  |  |  | distributed errors
| DUT Output Match* | 1.5 | VSWR | Fixed | Fixed | S22, ρ | 0.2 | (F12/F1)-(F2/F1G1)= | 0.99499
| Freq Converting DUT? | Non-conv'ing | (e.g. Amplifier, Attenuator) |  |  | F12 (Combined F) | 2.0853 |  |  | Factor for First ρ |  | Factor for Second ρ |  | Double-sided coverage factor, k
Noise Source | [5 empty] | NF12 (F12 in dB) | 3.1916 | Mismatch Standard Uncert, dB |  | Style Conv | Distrib Factor | Style Conv | Distrib Factor | 1 | 2 | 3
| ENR Uncertainty | 0.15 | dB | 95th %ile (2σ) | Gaussian |  |  | NS to DUT in | 0.035398 | 0.712 | 0.57776 | 1 | 1 | Percentile coverage from k
| Noise Source Match* | 1.15 | VSWR | Maximum | Rayleigh | Noise Source ρ | 0.069767 | NS to Instrument | 0.018746 | 0.712 | 0.57776 | 1 | 0.57776 | 68.27% | 95.45% | 99.73%
[8 empty] | DUT out to Instr | 0.13235 | 1 | 1 | 1 | 0.57776 | Single-sided k for the same %ile
Instrument | Instrument Noise Fig | 10 | dB | Fixed |  | F2 (inst), Power Ratio | 10 |  |  |  |  | Selected Input: |  | 0.47523 | 1.6901 | 2.7822
| Instrument Uncertainty for NF | 0.02 | dB | Maximum (3σ) | Gaussian |  |  | Combined Uncertainties, User Cal |  |  |  | 95th %ile | Rayleigh
| Instrument Uncertainty for Gain | 0.07 | dB | Maximum (3σ) | Gaussian |  |  | δNF12 | 0.036014 | dB |  |  |  | Order of Combination, N |  | 15
| Input Match* | 1.45 | VSWR | 95th %ile | Rayleigh | Instrument ρ | 0.18367 | δNF2 | 0.019891 | dB | δNF2, int cal | 0.26511 | dB
| NFE Improvement/Int Cal | 9 | dB | 95th %ile (2σ) | Gaussian |  |  | δG1 | 0.13993 | dB | δNF2, no cal | 3.0103 | dB | Formula term for combining
[8 empty] | δENR | 0.075 | dB |  |  |  | 1 | 1.9889 | 2.923
External Preamp | Configuration | Not used | (Items below applicable w/Ext Preamp only)
| Noise Figure, Preamp Alone | 3.5 | dB | Fixed |  | F2 (p.a. only), Pwr Ratio | 2.2387 | Sens * Uncertainties, User Cal |  |  | Relative Contribution (User Cal) |  |  | δNF2 * Sens, uncal |  | 0.15087
| Gain | 21 | dB | Fixed |  | G (preamp), Pwr Ratio | 125.8925 | δNF12 * F12/F1 | 0.037638 | dB | 20.16%
| Input Match* | 1.3 | VSWR | Fixed | Fixed | Ext Preamp ρ | 0.13043 | δNF2 * F2/F1G1 | 0.00099689 | dB | 0.01414%
| NFE Improvement/Int Cal | 9 | dB | 95th %ile (2σ) | Gaussian |  |  | δG1 * (F2-1)/F1G1 | 0.0063119 | dB | 0.567% |  | Summary:
[6 empty] | F2 for combined pre- | 2.3102 | δENR * ((F12/F1)… | 0.074624 | dB | 79.26% |  |  | Standard | Expanded | 3σ
| * Match terms may be entered as the return loss (RL) in dB (Sxx), or VSWR, or as a reflection |  |  |  |  | amp and instrument | [6 empty] | Calibration Used | (σ) | 2σ, ≈95% | (≈99.73%)
| coefficient, ρ. Example: RL = -15 dB is equivalent to VSWR = 1.43 is equivalent to ρ = 0.178. |  |  |  |  | F2 for selected config | 10 | RSS Combined | 0.08366 | dB |  |  | Normal (User Cal) | 0.08366 | 0.16732 | 0.25098 | dB
| For a perfect match, enter 1 for VSWR = 1; for a total reflection, enter 0 for RL = 0 dB. |  |  |  |  | (w/ or w/o Ext Preamp) |  | w/Internal Cal | 0.084482 | dB |  |  | Internal Cal | 0.084482 | 0.16896 | 0.25345 | dB
[8 empty] | w/o Cal, w/o δNF2 | 0.083654 | dB |  |  | Uncalibrated | 0.19063 | 0.29226 | 0.38364 | dB
Style:
Computed Uncertainty |  | 2σ: Expanded | Coverage | Level of | Note
| Calibration Used | Uncertainty | Factor, k | Confidence
| Normal (User Cal) | ±0.167 dB | 2 | ≈95%
| Internal Cal | ±0.169 dB | 2 | ≈95%
| Uncalibrated | ±0.292 dB | 2 | ≈95%
References | Noise Figure Measurement Accuracy--The Y-Factor Method; Keysight Application Note 57-2 |  |  | url: http://www.keysight.com/find/AN57-2
| Revisiting Mismatch Uncertainty with the Rayleigh Distribution |  |  | url: http://www.keysight.com/find/RayleighMismatch
| Guide to the expression of uncertainty in measurement (GUM) |  |  | url: http://www.bipm.org/utils/common/documents/jcgm/JCGM_100_2008_E.pdf
| Effects of Noise Source Mismatch on Y-Factor Noise Figure Measurements in the Millimeter-Wave Range |  |  | See the DraftOfNoiseSourceMismatch tab for a 2017 draft of this article
Tables zone
Distribution Factors
| Rayleigh | 0.57776
| Fixed | 1
| Uniform inside Circle | 0.70711
| GammaStyles | Annotation | RayConv-95th | WithinCircConv
| Maximum | ¦Γmax¦ | 0.712 | 1
| 95th %ile | ¦Γ95¦ | 1 | 1.026
| 80th %ile | ¦Γ80¦ | 1.269 | 1.118
| Median | ¦Γ50¦ | 2.079 | 1.4142
| Mean | ¦Γmean¦ | 1.953 | 1.5
| Fixed | ¦Γ¦ | 1 | 1
Amp/Recr
| Non-conv'ing | 1
| Yes | 0
| Uncertainty Style | k | Coverage (text)
| 1σ: Standard | 1
| 2σ: Expanded | 2 | ≈95%
| 3σ | 3 | 99.73%
External Preamp
| Not used | 0
| In use | 1
Instructions: Fill in information in column C (highlighted in yellow). In lightly highlighted cells in columns
E and F, additional information on the statistics of the values may be entered. Observe the results in blue.
Another uncertainty: Devices whose output noise is highly sensitive to changes in source impedance can have additional uncertainties in NF--these are not included in this spreadsheet. Such devices are best measured with low ENR noise sources specified to change reflection coefficient by less than 0.01 between states: Keysight 346A & N4000A.
Instructions: Fill in information in column C (highlighted in yellow). In lightly highlighted cells in columns
E and F, additional information on the statistics of the values may be entered. Observe the results in blue.
Another uncertainty: Devices whose output noise is highly sensitive to changes in source impedance can have additional uncertainties in NF--these are not included in this spreadsheet. Such devices are best measured with low ENR noise sources specified to change reflection coefficient by less than 0.01 between states: Keysight 346A & N4000A. See also reference 4 ("Effects of Noise Source Mismatch...") for mmW issues.
Note on A3: This calculator is significantly more complicated than the legacy NF Measurement Uncertainty Calculator from 1999. (That earlier calculator is included on a tab in this calculator.) What are the differences? This calculator adds two additional estimates of the uncertainty: the "without calibration" case and also with the "Internal Cal" feature of some analyzers. This calculator uses knowledge of the statistical distribution of error sources such as ENR Uncertainty (2∙sigma) and Instrument Uncertainty (3∙sigma) and combines those through the method of the GUM (ISO/IEC Guide to the expression of uncertainty in measurement). This calculator recognizes the statistics of mismatch in the analyzer and the noise source, and treats those as recently researched and reported in the reference on mismatch uncertainty--see the References rows on this sheet. This calculator has controls for external preamplification. This calculator will not get the same results as the simpler calculator because of the different assumptions. In the most critical DUTs, the ENR uncertainty usually dominates, and the calculators will agree for the case combination of User Cal and 95% confidence specifications. When mismatch errors are significant, this calculator will be more accurate, though harder to use because of all the fields in columns E and F that can be filled in. The default settings of those fields will be optimum for most applications.
Note on C3: Another (on-line) version of Uncertainty Calculator is available on the Keysight website. That version has unique features. It has default data for Keysight noise sources and analyzers inside the calculator. It has the ability to sweep a parameter, such as DUT gain, and observe the effect of that parameter on uncertainty. This version has unique features of exposing all equations to the user and showing the statistical basis of some of the mathematics. Agilent/Keysight ESA, PSA and X-Series analyzers with Noise Figure Measurement applications have a version of Uncertainty Calculator built-in under the Mode Setup key.
Note on E4: Specifications policies can be highly varying. In this column, the places where the policies can vary are the "match" rows. Select 95th percentile or median if the specification is given or known in one of those forms. "Maximum" is used for specifications given as warranted values. 80th percentile is useful for "typical" specifications when "typical" is defined as 80th percentile, such as in some Keysight products. In some circles, "Typical" is the Mean. In the rows in which the style is not settable, the statistical treatment used is shown for information only. For example, ENR Uncertainty is given in Calibration Reports for Keysight Noise Sources as a 95th percentile, with a coverage factor k = 2. The two "Instrument Uncertainty" values are known to be better than 3∙sigma performance in Keysight X-Series analyzers; Keysight recommends treating NFA specifications as 3∙sigma as well.
Note on F4: The statistical distribution of the specification is either settable (in the cases shown in blue) or given for information only in this column. The cases in which it is settable are all match cases. "Fixed" means that the match is always at the given level. "Rayleigh" means that the probability density function (pdf) of the magnitude of the reflection coefficient, ρ, is a Rayleigh distribution. This is known to be the case with Keysight X-Series spectrum analyzers, as shown in their Specfications Guide and in the reference on Rayleigh mismatches. "Uniform within a circle" is a distribution that was historically used in the industry, before the Rayleigh research, as documented in Keysight Application Note 1449-3, but is rarely accurate.
Note on A10: A snapshot of the specifications of all Keysight Noise Sources at the time of release of this spreadsheet is included in the tab "NoiseSourceSpecs."
Note on C11: There are two ways to estimate the ENR uncertainty: the calibration report for the noise source in use, at the frequency of interest; or from the noise source Operating and Service Manual, Table 1-1. These two sources of information are described in detail on the CoverageFactorDocumentation tab. The calibration report gives the more accurate estimation of uncertainty. Also, in frequency converting DUT applications and other two-noise-source applications, the ENR uncertainty and even the Noise Source Match can be different for the two locations of usage of the noise source or noise sources. This calculator is not designed to handle that case ideally; we recommend you use the worse of the parameters of the two applications to get a conservative, instead of accurate, estimate of noise figure uncertainty.
Note on C12: One way to fill in this field is to use the match information for the selected noise source as found in its Operating and Service Manual, Table 1-1. This specification for Keysight Noise Sources is a warranted (maximum) value. Observations of the statistics of this match do not show that it is Rayleigh-distributed. But those observations do show that the standard uncertainty computed by assuming that the distribution is a Rayleigh pdf is either an accurate or conservative (depending on model) estimate of the uncertainty. Using the Rayleigh distribution and "Maximum" always results in a more accurate model than previous practice or the other distributions. Therefore, a recommended (and default) setting of the "Specification Style" column is Maximum and of "Distribution" is Rayleigh for Keysight Noise Sources. A second way to fill in this field is from the calibration report of the actual noise source being used at the frequency of interest. In this case, change the Specification Style to Fixed and the Distribution column also to Fixed.
Note on A14: A snapshot of the specifications of the Keysight NFA (Noise Figure Analyzer) and X-Series Signal Analyzers with NF measurement capability is included in the tab "AnalyzerSpecs." These specifications are the most recent ones available at the time of the release of this spreadsheet.
Note on C14: The NF of the analyzer is highly frequency dependent. Graphs of NF versus frequency are available in the Specifications Guides for Keysight X-Series analyzers, in the Noise Figure Measurement chapter. There is also information there on how to compute NF from the common spectrum analyzer specification of DANL (Displayed Average Noise Level).
Note on C15: For Keysight X-Series Analyzers, this value comes from the Specifications Guide, Noise Figure Measurement Chapter, Noise Figure table; the column labeled "Instrument Uncertainty."
Note on C16: For Keysight X-Series Analyzers, this value comes from the Specifications Guide, Noise Figure Measurement Chapter, Gain table; the line labeled "Instrument Uncertainty."
Note on C17: Information on Input Match of Keysight analyzers is shown in their Specifications Guides. You will see that EXA, MXA and PXA have a reflection coefficient that is very well modeled with a Rayleigh distribution. The 95th percentile VSWR for each RF mixing band with the internal preamp and 0 dB attenuation is tabulated for MXA and PXA. For NFA and CXA, the frequency dependence of mismatch is highly reproducible but not very well modeled as Rayleigh distributed. In that case, Keysight recommends reading that VSWR from a graph and using "Fixed" for the Style and Distribution columns.
Note on C18: NFE Improvement is a specification of the apparent reduction in the noise floor of an analyzer using Noise Floor Extension (NFE). The 95th percentile improvement is documented in the Specifications Guide for Keysight analyzers that have this feature. As of this writing, this feature is standard in PXA and UXA models, and optional in MXA and EXA. The use of the models from NFE in the Noise Figure Measurement Application is called "Internal Cal" in analyzers that have NFE. The value entered here is irrelevant for analyzers without NFE, because it only affects the "Internal Cal" result. The only analyzers with Internal Cal are those with NFE. Leaving the default setting in this cell is appropriate for analyzers without NFE.
Note on A20: Joe Gorin: A snapshot of the specifications of the Keysight USB preamplifiers, U7227A/C/F, is included in the tab "ExtPreampSpecs." These specifications are the most recent ones available at the time of the release of this spreadsheet.
Note on C24: NFE Improvement is a specification of the apparent reduction in the noise floor of an analyzer/external preamp system using Noise Floor Extension (NFE). The 95th percentile improvement is documented in the specifications of the preamp. As of this writing, there are Keysight preamps with this feature but there are no supported specifications on this feature, though the default value of 9 dB is a good estimate. The use of the models from NFE in the Noise Figure Measurement Application is called "Internal Cal" in analyzers that can use preamplifiers with NFE noise characterization tables. The value entered here is irrelevant for the cases where either the analyzer or the preamp does not have NFE, because it only affects the "Internal Cal" result. The only analyzer/preamp combinations with Internal Cal are those with NFE. Leaving the default setting in this cell is appropriate for analyzer/preamp combinations without NFE.
Note on C31: "Expanded" uncertainty, sometimes called 95th percentile, is an industry-standard form for the expression of measurement uncertainty.
Note on F31: There are even more sources of uncertainty in addition to those listed and the one in the text block below this cell. These additional ones often are negligible or can be minimized. They are listed here to improve completeness. * Measurement repeatability. This can be observed experimentally. If desired, it can be minimized with increased averaging. * Reverse isolation. The model assumes infinite reverse isolation. So the uncertainty might be different for devices (example: passive devices) for which this is not the case. * Cold Temperature Estimate. The estimate of the noise-source-off (cold) temperature may be in error. This temperature is sensed with errors near one degree Kelvin in the N4000 series of noise sources, but is only estimated by the user for the 346 series. The thermal conductivity of the package of the 346 series is such that the temperature is more closely related to the DUT chassis temperature than to the ambient temperature; this can make manual estimation error-prone. * Cold Temperature Absolute. Noise source ENRs are only specified over a range of +/-1 degree Kelvin; the typical temperature coefficient is under 0.01 dB/K change in ENR with temperature. See the tab, "TCofENR."
Note on B34: Internal Cal is standard on the UXA and PXA, and is Option NFE on the EXA and MXA.
Note on D71: Agilent 2012: Conversion factor from the the given style to the "reference" style (which is 95th percentile) for Rayleigh distributions
Note on E71: Agilent 2012: Conversion factor from the the given style to the "reference" style (which is "Maximum") for uniform-within-Circle distributions

## AnalyzerSpecs
This sheet shows the specifications for Keysight analyzers with NF measurement capability.
Please refer to the most recent version of the Specifications Guide for the most up-to-date information.
Click to jump to model:
CXA | EXA | MXA | PXA | UXA | X-Series
ESA | PSA |  |  |  | Older analyzers
NFA | NFA-B |  |  |  | Dedicated Instruments
CXA | From document N9000-90016, 01-Mar-2012
CXA | Instrument Noise Figure
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA | Instrument Uncertainty for Noise Figure
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA | Instrument Uncertainty for Gain
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA | VSWR
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
CXA
ESA | From document E4401-90490, January 2011
ESA | Instrument Noise Figure
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA | Instrument Noise Figure can also be computed from DANL (Displayed average noise floor) as stated in the footnotes of the Specifications Guide;
ESA | Nominally, NF = D - (K - L + N + B)
ESA | where D = DANL (displayed average noise level),
ESA | K is kTB (-173.98 dBm in a 1 Hz bandwidth at 290 K)
ESA | L is 2.51 dB (the effect of log averaging used in DANL verification)
ESA | N is 0.52 dB (the ratio of the noise bandwidth of the RBW filter with which DANL is
ESA | specified to an ideal noise bandwidth, for the ESA with its 4-pole synchronously
ESA | tuned RBW filters)
ESA | B is ten times the base-10 logarithm of the RBW (in Hz) in which the DANL is
ESA | speciffied. B is 30 dB for the 1 kHz RBW
ESA | The actual NF will vary from the nominal due to frequency response errors
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA | E4402 with Preamp (Option 1DS)
ESA
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA | E4404/5/7 with Preamp (Option 1DS)
ESA
ESA | Instrument Uncertainty for Noise Figure
ESA | NOTE: The statistical distribution of IU for NF in ESA is unusual, and not Gaussian. Therefore, the calculator does not make an accurate estimation
ESA | of NF uncertainty for ESA, because it does not account correctly for the distribution. Worst case addition of uncertainties might work best for ESA.
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA | Instrument Uncertainty for Gain
ESA | NOTE: The statistical distribution of IU for Gain in ESA is unusual, and not Gaussian. Therefore, the calculator does not make an accurate estimation
ESA | of NF uncertainty for ESA, because it does not account correctly for the distribution. Worst case addition of uncertainties might work best for ESA.
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA | VSWR
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
ESA
EXA | From document N9010-90025, 01-Sept-2015
EXA | Instrument Noise Figure, Options P03, P07, P13 and P26
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | Instrument Noise Figure can be computed from DANL (Displayed average noise floor) as stated in the footnotes of the Specifications Guide;
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | Instrument Uncertainty for Noise Figure
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | Instrument Uncertainty for Gain
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | VSWR
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | VSWR, 3.6 to 26.5 GHz, Option P07, P13 or P26; this data is from PXA but is representative
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA | VSWR, tabular version, P03, P07, P13, P26; see footnote a for P32 and P44
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
EXA
MXA | From document N9020-90113, 01-Feb-2012; VSWR table from 01-Jan-2017
MXA | Instrument Noise Figure
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA | Instrument Uncertainty for Noise Figure
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA | Instrument Uncertainty for Gain
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA | VSWR, graphical version
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA | VSWR, tabular format
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
MXA
NFA | From document 5980-0164E, 2007-11-05
NFA | Instrument Noise Figure
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA | Instrument Noise Figure, textual version:
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA | Instrument Uncertainty for Noise Figure and Gain
NFA
NFA
NFA | Instrument Uncertainty for Noise Figure and Gain
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA | VSWR
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA
NFA-B | The NFA-B models (N8973B, N8974B, N8975B, N8976B)
NFA-B | NFA datasheet:https://www.keysight.com/us/en/assets/9018-04730/technical-specifications/9018-04730.pdf
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B | Instrument Noise Figure can also be computed from DANL (Displayed average noise floor) as stated in the footnotes of the Specifications Guide;
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B | DANL with Internal Preamp off
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B | Instrument Noise Figure Uncertainty and Instrument Gain Uncertainty
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B | DANL with Internal Preamp On
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B | Instrument Noise Figure Uncertainty and Instrument Gain Uncertainty
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
NFA-B
PSA | From document E4440-90647, 01-Dec-2012
PSA | Instrument Noise Figure
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA | Option 1DS, 0 to 3 GHz
PSA
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA | Option 1DR, 0 to 3 GHz
PSA
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA | Option 110, 3 to 50 GHz
PSA
PSA | Instrument Uncertainty for Noise Figure
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA | Instrument Uncertainty for Gain
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA | VSWR, graphical format
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA
PSA | VSWR, tabular format
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA | E4440/43/45
PSA
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PSA | E4446/47/48
PXA | From document N9030-90017, 01-Jan-2017
PXA | Instrument Noise Figure, Graphical Version
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA | Preamp option ≤ P26 |  | i.e. max Freq ≤ 26.5 GHz
PXA
PXA | Instrument Noise Figure can also be computed from DANL (Displayed average noise floor) as stated in the footnotes of the Specifications Guide;
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | Instrument Noise Figure, Tabular Version
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | Instrument Uncertainty for Noise Figure
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | Instrument Uncertainty for Gain
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | VSWR, graphical version
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | VSWR, tabular format
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA | NFE Improvement (Internal Cal)
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
PXA
UXA | From document N9040-90002, 01-Mar-2017
UXA | Instrument Noise Figure can be computed from DANL (Displayed average noise floor) as stated in the footnotes of the Specifications Guide;
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA | Instrument Uncertainty for Noise Figure
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA | Instrument Uncertainty for Gain
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA | VSWR, graphical version
UXA
UXA | VSWR, graphical version
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA | VSWR, tabular format
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA | NFE Improvement (Internal Cal)
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA
UXA

## NoiseSourceSpecs
This sheet shows the specifications for Keysight noise sources.
Please refer to the most recent version of the Operating & Service Manual for the most up-to-date information on the population of noise sources for each model.
Please refer to the most recent Calibration Report for the individual noise source being used for the most accurate information.
Click to jump to model:
346A/B/C | 346C-K01 | 346C-K40 | R347 | Q347 | N400xA
346A/B/C | From document 000346-90148, 26-May-2017
346A/B/C | VSWR
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C | ENR Uncertainty
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346A/B/C
346C-K01 | From document 00346_90142, 01-Mar-2015
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K01
346C-K40 | From document 00346_90153, 01-Feb-2016
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
346C-K40
R347B | From document 00347-90013, August 2014
R347B | VSWR
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B | ENR Uncertainty
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
R347B
Q347B | From document 00347-90013, August 2014
Q347B | VSWR
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B | ENR Uncertainty
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
Q347B
N400xA | From N4000A, N4001A, N4002A SNS Series Noise Sources - Technical Overview, Literature Number 5988-0081EN, January 27, 2016
N400xA | VSWR, tabular version
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA | VSWR, graphical version
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA | ENR Uncertainty, example
N400xA | (For best accuracy of estimated uncertainty, use the calibration report from the particular device used.)
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA
N400xA | Field calibrations will change these uncertainties
N400xA
N400xA | At the time of this writing, September, 2017, Keysight does not have a "Standards Lab" field calibration available
N400xA | for the N4000 series. As a result, field calibration will use a reference that is two steps removed from the primary reference,
N400xA | making the calibration "against a twice removed reference" and inferior to that of a new N4000 series device. This effect is
N400xA | approximately an increase in ENR Uncertainty of 0.015 dB. Better performance can be achieved by replacing N4000 noise
N400xA | sources every two years with new devices instead of recalibrating them every two years.
N400xA

## ExtPreampSpecs
This sheet shows the specifications for the Keysight USB Preamplifier family, U7227A/C/F
These were designed specifically for NF use.
Click to jump to model:
U7227A | 0.01 to 4 GHz
U7227C | 0.1 to 26.5 GHz
U7227F | 2 to 50 GHz
All data from document 5991-4246EN, 24-Apr-2014
U7227A | Gain
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A | Noise Figure
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A | Return Loss
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227A
U7227C | Gain
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C | Noise Figure
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C | Return Loss
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227C
U7227F | Gain
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F | Noise Figure
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F | Return Loss
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F
U7227F

## LegacyCalc
© Agilent Technologies, Inc. 1999 |  |  | This sheet is the simpler but less complete version of NFUC from 1999
| dB | Linear
DUT NF, F1= | 3 | 1.9953 |  | F12/F1= | 1.0451
Instrument NF, F2= | 10 | 10 |  | F2/F1G1= | 0.050119
DUT GAIN, G1= | 20 | 100 |  | (F2-1)/F1G1= | 0.045107
Combined NF, F12= | 3.1916 | 2.0853 |  | (F12/F1)-(F2/F1G1)= | 0.99499
Match | Units* | Refl Coef |  |  | Negative | Positive | Max
Noise Source= | 1.1 | 0.047619 |  | Uncert NS-DUT IN= | 0.083119 | 0.082331 | 0.083119
DUT Input= | 1.5 | 0.2 |  | Uncert NS-NFA= | 0.11899 | 0.11738 | 0.11899
DUT Output= | 1.5 | 0.2 |  | Uncert DUT OUT-NFA= | 0.51108 | 0.48267 | 0.51108
Instrument= | 1.8 | 0.28571
Uncertainties | dB
Instrument NF= | 0.05 |  |  | Uncert NF12= | 0.096748
Instrument Gain= | 0.15 |  |  | Uncert NF2= | 0.12877
Noise Source ENR= | 0.1 | (Amplifiers Only) |  | Uncert G1= | 0.54642
Noise Source ENR= | 0 | (Receivers Only) |  | Uncert ENR= | 0.1
|  |  | Noise Figure Uncertainty = |  | 0.14337 | dB
* This term can be entered in dB (Sxx), VSWR or as a reflection coefficient. | [5 empty] | D.Boyd 1999
e.g. -15dB = 1.43 VSWR = 0.178 reflection coefficient
This spreadsheet calculates the uncertainty of noise figure measurements. The numbers in yellow are user variables and the green area is a calculation area. The final uncertainty is shown in blue, this should be added to and subtracted from the result shown on the noise figure measurement instrument in order to give the spread of possible values.

## RevHist
Revision History for calculator replacing the Legacy version of 1999-2002
Revised: | 2012-10-18
Content: | New creation, leveraged from Noise Figure Uncertainty Calculator of 1999-2002
Also leveraged from Agilent Average Power Sensor Uncertainty Calculator, 2008-2011
Revised: | 2015-02-02
Content: | Add UXA and U7227 External Preamp to Specs tabs
Change Agilent to Keysight
Add ESA and PSA to Specs tabs
Change references from hyperlinks to copy-and-paste urls to avoid security concerns
Add navigation links in Specs tabs
Revised: | 2015-03-16
Content: | Comments updated
Revised: | 2015-05-29
Content: | Comments updated
Revised: | 2016-03-04
Content: | Added NFA-B information. Typographical errors fixed in comments, including ESA specifications.
Also: 2016 will probably see the completion of a paper on how noise source mismatch
causes further errors that cannot accurately be captured by this calculator. The working title of the paper is
"Effects of Noise Source Mismatch on Y-Factor Noise Figure Measurements in the Millimeter-Wave Range"
Revised: | 2016-10-07
Content: | Updated EXA, including adding mmW model
Added a draft of the new reference described in the paragraph above
Revised: | 2016-12-06
Content: | Added tab of TC of ENR, and discussion in the comment on Calculator sheet cell F31
Updated N400xA Noise Source information
Revised: | 2017-06-09
Content: | Updated the DraftOfNoiseSourceMismatch tab with a new copy of the paper. Made it an icon to
be double-clicked instead of "picture" of the pages.
Revised: | 2017-09-19
Content: | Updated the DraftOfNoiseSourceMismatch tab with a new copy of the paper.
Added 346C Option K01 and 346C Option K40 noise sources to the NoiseSourceSpecs tab
Updated other items on the NoiseSourceSpecs tab
Updates items on the AnalyzerSpecs tab
Revised: | 2024-05-09
Content: | Updated the NFA-B specs under AnalyzerSpecs Tab

## TCofENR
Keysight noise sources achieve their low uncertainty through calibration against primary noise sources (hot and cold resistors) at an NMI (National Metrology Institute),
typically NPL (National Physical Laboratory of the UK). A device calibrated against this primary reference then is used as a secondary reference to calibrate another
noise source. For the Keysight N4000 series, this noise source is the production noise source shipped to the purchaser; for the 346 series, this noise source is
used as a reference for production uses. In that sense, the N4000 series is calibrated against a one-level removed (from primary) noise source, and the 346 series
is calibrated against a two-level removed noise source.
If the TC of ENR is consistent in the product line, the temperature at which any secondary calibrations occur is irrelevant, because the effects cancel in the reference
and the DUT. That is a resonable approximation. Here is a graph of the observed TC of the N4002A noise source:
Over most of the frequency range, the TC is very close to that of the TC of the avalanche diode voltage, so we would expect the whole product line to be consistent with
that fundamental property of the noise diode.
But the calibration at NPL is highly relevant. The calibration of the product line will vary with the temperature at which NPL does their calibration. Although NPL maintains
their calibration environment to a +/-1 degree K tolerance around 23 degrees C (296.15 K), the ambient temperature is only indirectly of importance. The temperature
of the noise diode depends on the thermal environment of the noise source, which is likely dominated by the equipment to which it is connected by its RF connector, and
it depends on the internal rise in the noise source due to its internal power dissipation.
Typically, the NPL environment makes for an internal temperature of 302.8 K for 18 GHz devices (346A/B, N4000/1A) and 304.8 K for 26.5 GHz devices (346C, N4002A).
To the extent that a usage environment causes variations from these internal temperatures, there will be a bias in ENR according to this TC. At most frequencies, this
can be a modest error. For a +3 K change, the ENR will nominally increase by +0.013 dB.

## CoverageFactorDoc
This page will document the "coverage factors." The coverage factor is the relationship
between the standard deviation of a parameter and its specification. This page will show
the relationships for three parameters:
1 | ENR Uncertainty (Coverage factor k = 2)
2 | Instrument Uncertainties (Coverage factor k = 3)
3 | Noise Source match (can be treated as "Max"/"Rayleigh")
ENR Uncertainty, with coverage factor k = 2
The ideal place to find the ENR uncertainty is in the calibration report for the noise source being used. Here is an example of page 3 of the report:
The third column shows the "ENR UNC," the uncertainty in ENR which is needed in this spreadsheet.
Note the sentence at the bottom: "Uncertainties are calculated using a coverage factor of 2." These, then, are
95th percentile "expanded" uncertainties. This spreadsheet divides them by two to find the "standard
uncertainty," which is the standard deviation of the distribution.
Another possible source of data for ENR uncertainty is from the Operating and Service manual of the
noise source, such as literature number 00346-90148. Here is a copy of part of one page from that manual:
These numbers are larger than those shown in the example calibration report.
Note the statement at the bottom of this page: 90% of instruments with 90% confidence. What this is saying is that for a
sample population of noise sources, the calibration report for more than 90% of them were better than the numbers
tabulated. (How many more than 90% depends on the sample size used, in order to assure 90% confidence in the
90% coverage value.) Thus, you are somewhat "safe" in using the values from the Operating and Service Manual,
but the numbers from the calibration report are usually smaller and always a more accurate estimate of the uncertainty.
Instrument Uncertainties, with coverage factor k = 3
There are two "instrument uncertainties" used in noise figure uncertainty analysis: One for NF and one for gain. These
are uncertainties due to nonlinearities in the analyzer response to noise over the range of levels that can be excited
during the measurement process.
The NF instrument uncertainty is the nonlinearity across a range of the possibly seen levels between switching the the noise
source on and off. Therefore, it can depend on the ENR of the noise source. The "Gain" instrument uncertainty is the
nonlinearity across a range excited by the gain of the DUT, as experienced between the "cal" and "measure" parts of
the measurement process, a potentially much wider range.
Here is a statistical plot of the Instrument Uncertainty for NF observations from MXAs:
The three x-values correspond to the maximum ranges using three different noise source series with 6.5, 17 and 22 dB
maximum ENRs. The y values shown are the maximum absolute values of the nonlinearities across 6.5, 17 and 22 dB
ranges anywhere within a wide range of noise levels. Because these are the maximum values, the distribution cannot be
expected to be Gaussian, and the levels shown are higher than those experienced in most measurements.
The error at any particular noise level should be Gaussian according to the central limit theorem. Because we are
showing the maximum error in this plot, and that level is well under the warranted levels of 0.020, 0.025 and 0.030 dB
for 6.5, 17 and 22 dB ENRs respectively, we can be confident that treating these warranted specifications as 3σ
values is appropriate. Thus, the coverage factor k = 3 for this parameter.
The available statistics for the Instrument Uncertainty for Gain are in a different format:
This specification is set to 0.07 dB to ensure continuing producibility even with process drifts, though that is obviously
very conservative. In actuality, the Instrument Uncertainties are so small that measurement errors dominate our
observations of them.
Noise Source match, Statistical Treatment as "Max"/"Rayleigh"
The mismatch uncertainty at the interface between the noise source and the DUT during measurement, and between
the noise source and the analyzer input during user calibration, is an error contributor used in this spreadsheet.
Research in 2011 has shown conclusively (see the reference on the Calculator sheet) that treating the statistical distribution
of the mismatch as a Rayleigh distribution of magnitude is an excellent way to accurately compute errors for many
devices. Examples are shown in that reference for signal analyzers (such as would be used for noise figure analysis),
power sensors and signal sources.
Analysis shows that the Rayleigh model is not accurate for noise sources. However, using the Rayleigh model with
the match set to the warranted maximum magnitude for the noise source gives either accurate or conservative results.
Therefore, that is the recommended practice for using this spreadsheet.
Here is an example: the magnitude of the reflection coefficient for 20 346B noise sources with the noise source on.
In uncertainty analysis, the parameter of interest is the "standard uncertainty." This is the RMS error. For mismatch
errors, the mismatch standard uncertainty is proportional to the RMS value of the reflection coefficient, ρ.
In the example illustrated, the RMS value of ρ for the 0.3 to 5 GHz case is 0.0226.
The stated maximum ρ for this frequency range for this noise source is 0.07.
If we assume a Rayeigh distribution and a maximum (3σ) ρ of 0.07, this is equivalent to an RMS ρ that is 28% larger
than the actual RMS ρ.
Similarly, using a maximum ρ of 0.11 for the 5 to 18 GHz case gives a higher RMS ρ than the observed ρ by 40%.
Therefore, one convenient way of configuring this Uncertainty Calculator is by entering the maximum ρ from the
Table 1-1 specifications from the Installation and Service Manual of the noise source in the "Noise Source match"
cell of the calculator sheet, and selecting "Maximum" for the "Specification Style" on the same row, and "Rayleigh"
for the distribution on that row.

## DraftOfNoiseSourceMismatch
Draft of Noise Source Mismatch article.
Double click on the icon to open it.