## Special Calibration Matrix
Special Calibration Matrix |  |  | ETE Model / Special Calibration Option Required
Platform | N7800A Software Application --- Keysight Calibration Application Number and Name | Footnotes | 11667A | 11667B | 11667C | 8491A/B-003 | 8491A/B-006 | 8491A/B-010 | E4419A/B Ref Port | 8494G/H | 8496G/H | 8481A | 8481D | 8482A | 8485A | 8485D | 8487A | 8487D | E9300A | E9300A-H25 | E9301A | E9304A | E9304A - H18 | E9304-H19 | E9304-H20 | V8486A+V281A | N8481A - Std | N8481A - CFT | N8481H | N8482A - Std | N8482A - CFT | N8482H | N8485A - Std | N8485A - CFT | N8485A Opt 033 | N8486AR | N8486AQ | N8487A - Std | N8487A - CFT | N4691BH33
E6601A | N7805A E6601A | 1,29,30 | [16 empty] | KC+U |  | KC+U | KC+U
PSA | N7810A PSA Series | 1,13,15,22,29 |  |  |  | H33 | H33 | H33 |  | H50 | H50 | H84 |  | H84
ESA | N7811A ESA Series | 1,15,22 |  |  |  |  | H33 | H33 |  | H50 | H50 |  | H84
E4406A | N7812A E4406A | 1,12,22 |  |  |  |  | H33 | H33 |  | H50 | H50 |  | H84 | H84
CSA | N7813A CSA Series | 1,14,15,20 |  |  |  |  | H33 |  |  | H50 | H50
X-Series | N7814A X-Series Signal Analyzer | 1,12,15,19,29 |  |  |  | H33 | H33 | H33 |  | H50 | H50 |  |  | H84 | KC+U | KC+U | [6 empty] | KC+U | [6 empty] | KC+U | KC+U |  | KC+U | KC+U
896XX VSA | N7815A 896XX VSA Series | 1,12,15,19,29 |  |  |  |  | H33 |  |  | H50 | H50 | KC+U |  | KC+U
EMI | N7817A EMI Receiver | 1,13,15,22,29 |  |  |  |  | H33 | H33 |  | H50 | H50 | H84 |  | H84 | [7 empty] | KC+U
X-Series EMI | N7818A MXE EMI Receiver | 33 |  |  |  | H33 | H33 | H33 | [6 empty] | KC+U | [7 empty] | KC+U | [9 empty] | KC+U | KC+U
PSG-D | N7820A PSG Series | 1,10,11,26,29 | [11 empty] | H84 | H84 | H84 | H84 | H84 |  | KC+U |  | KC+U | KC+U | KC+U |  | H67 |  | KC+U |  | KC+U | KC+U |  | KC+U | KC+U |  |  |  | KC+U | KC+U
ESG-B | N7821A ESG C Series | 1,4,29 | [11 empty] | H84 | [7 empty] | KC+U
MXG | N7822A MXG Series | 1,6, 7,29 | [14 empty] | H84 |  |  | KC+U |  | KC+U |  | KC+U
E8663B | N7823A E8663x Series | 1,3,29, | [17 empty] | KC+U |  | KC+U |  | KC+U
ESG-C | N7824A ESG B Series | 1,5,29 | [11 empty] | H84 | [7 empty] | KC+U |  | KC+U
PXB | N7825A PXB Series
PSG-C | N7826A PSG Series | 1,10,11,29 | [11 empty] | H84 | H84 | H84 | H84 | H84 |  | KC+U |  | KC+U | KC+U | KC+U |  |  |  | KC+U |  | KC+U | KC+U |  | KC+U | KC+U |  |  |  | KC+U | KC+U
NFA | N7830A NFA Series | 16,22,26 |  |  |  |  | H33 | H33
N4010A | N7831A N4010A | 1,8 | [9 empty] | H84 | H84
P-Series Power Meters | N7832A P-Series Power Meter | 1,21 | [6 empty] | Footnote 21 |  |  | KC+U | H84 | KC+U | [15 empty] | KC+U
N8300A | N7833A N8300A | 1,9,29 | [9 empty] | KC+U |  | KC+U | KC+U | [5 empty] | KC+U | KC+U
E6651A | N7834A E6651A | 1,2, 29 | [9 empty] | KC+U | [6 empty] | KC+U |  | KC+U | KC+U
PNA-L/X | N7840A PNA Series | 26,29 | [9 empty] | KC+U |  | KC+U | KC+U |  | KC+U | [8 empty] | H67 |  | KC+U |  | KC+U | KC+U |  | KC+U | KC+U |  |  |  | KC+U | KC+U
FieldFox RF Analyzer | N7841A FieldFox RF Analyzer | 1,15,22,29,32 |  |  |  |  | H33 | H33 |  | H50 | H50 | [11 empty] | KC+U |  | KC+U |  |  |  |  | H84
ENA | N7842A E5071C | 1,29,31 | [8 empty] | H52 |  |  | KC+U | [7 empty] | KC+U
PNA | N7843A Legacy PNA Series | 29 | [9 empty] | KC+U |  | KC+U | KC+U |  | KC+U | [10 empty] | KC+U |  | KC+U | KC+U |  | KC+U | KC+U |  |  |  | KC+U | KC+U
Upconverter | N7845A Upconverter Series | 1,18,29 | [11 empty] | H84 | H84 | H84 | H84 | H84 |  | KC+U |  | KC+U |  | KC+U
Downconverter | N7846A Downconverter | 1,13,15,22 |  |  |  | H33 | H33 | H33 |  | H50 | H50 | H84 |  | H84 | H84
Power Sensors | N7850A Power Sensors | 1,17,20,28,29 | H85 | H85 / H87 | H85 | [17 empty] | KC+U | [16 empty] | Footnote 20
U2000-Series Power Sensors | N7851A U2000 Series USB Power Sensors | 1,17,29 | H85 | H85 | H85 | [16 empty] | KC+U |  |  |  |  | KC+U |  | KC+U |  |  |  | KC+U
N5532A Power Sensor | N7852A N5532X Series Sensor Module | 23,25 | H86 | H86 | H86 | [6 empty] | Footnote 25 |  | Footnote 25 | Footnote 25 |  | Footnote 25 | [9 empty] | Footnote 25 | Footnote 25 |  | Footnote 25 | Footnote 25 |  | Footnote 25 | Footnote 25 |  |  |  | Footnote 25 | Footnote 25
USB Modular SMU | N7855A Source Measure Unit | 29
USB Modular Func. Generator | N7856A Function/Arbitrary Waveform Generator | 1,29 | [11 empty] | KC+U | [12 empty] | KC+U
Infiniium | N7860A Infiniium 90000 Series | 1,29 | [16 empty] | KC+U | KC+U
PXI Signal Generator | N7870A PXI Signal Generator | 1,29 | [19 empty] | KC+U | KC+U |  | KC+U |  |  |  |  | KC+U
PXI IF Digitizers | N7871A IF Digitizer | 1,29 | [12 empty] | KC+U | [17 empty] | KC+U | KC+U
PXI Vector Signal Analyzer | N7872A PXI Vector Signal Analyzer | 1,29 | [20 empty] | KC+U |  | KC+U
PXI Reference | N7873A PXI Frequency Reference | 1,29 | [11 empty] | KC+U | [7 empty] | KC+U | KC+U
PXI Vector Network Analyzer | N7874A PXI Vector Network Analyzer | 1,29 | [30 empty] | KC+U
CXA-m PXIe Signal Analyzer | N7875A CXA-m PXIe Signal Analyzer | 1,29 |  |  |  |  | H33 | H33 |  | H50 | H50 |  |  | H84 | KC+U | KC+U | [6 empty] | KC+U |  | KC+U |  |  |  |  | KC+U | KC+U |  | KC+U | KC+U
PXI Digital IO | N7876A PXI Digital IO | 1,29
PXI Modular Synthesizer | N7877A PXI Modular Synthesizer | 1,29 | [22 empty] | KC+U |  |  |  |  | KC+U
PXI Multiport Vector Network Analyzer | N7878A PXI Multiport Vector Network Analyzer | 1,29 | [11 empty] | KC+U | [8 empty] | KC+U | [6 empty] | KC+U
PXI Accessories | N7879A PXI Accessories | 1,29
Last Update - Apr, 4th, 2022
Note: "KC+U" is Keysight Calibration + Uncertainity
Footnotes: | Comments
1 | H99 - Calibration data on a CD in .csv format. (H99 is ONLY FOR POWER SENSORS) Note: The following models are not currenly supported for option H99. (Calibration data on a CD in .csv format.): N8486AQ, N8486AR, E9304A H20
2 | 8481A , E9300A , E9301A , E9304A
3 | E9304A , E9304A-H19 , E9300 H25
4 | 8482A H84, E9304A (alternate is E9304A-H19 )
5 | 8482A H84, E9304A
6 | MXG Power Sensors for N5161A/N5162A/N5181A/82A:
E9304A [use with standard model MXGs]
E9304A-H19 [use with opt 1EA MXGs]
7 | MXG Power Sensors for N5183A:
E9300A-H25 , E9304A H19 (for N5183-520)
8487A H84, E9304A-H19 (for N5183A-532 or 540)
8 | 8481D H84, 8481A H84
9 | 8481A , 8482A , 8485A , E9301A , E9304A
10 | PSG-C Power Sensors:
E9304A (alternate is E9304A-H19 )
E9300A-H25 (only for opt 520 models)
8487A H84 (alternate is 8485A H84 for opt 520 models)
8487D H84 (not required for Opt 520 model)
11 | PSG-D Power Sensors:
E9304A (alternates are E9304A H19 or 8482A H84)
E9300A H25 (only for opt 520 models)
8487A H84 (alternate is 8485A H84 for opt 520 models)
8487D H84 (not required for Opt 520 model)
V8486A+V281A
12 | 8482A H84 (8482A STD Note: The use of the standard model will
increase the measurement uncertainty. )
13 | 8482A H84 (8482A STD Note: The use of the standard model will
increase the measurement uncertainty. )
14 | Only supports 8494/6 H model.
15 | 8494G/H - H50, 8496G/H - H50
16 | 8494G/H - A Standard Keysight calibration is required.
17 | 11667A H85, 11667B H85, 11667C H85
18 | E9304A , E9304A-H19 , E9300-H25 , 8485A H84, 8487A H84
19 | 8482A H84
20 | N4691BH33 - Ecal Module for N8485-033 Power Sensor Calibration
21 | E4419A/B: The reference power port of the power meter must be calibrated, traceable to a national standard, to at least ±0.25% accuracy. It is assumed that this represents an end-of-period accuracy. Additionally, it is assumed that this represents a 95% confidence (k=2) coverage factor. It is necessary for the laboratory that executes this procedure to maintain this level of end-of-period accuracy. The calibrated reference accuracy drift rate is expected to be 0.005% per 6 months. This level of accuracy may only be achievable by a national metrology institute (NMI), such as National Physical Laboratory (NPL) ( http://www.npl.co.uk/) in the United Kingdom, or National Institute of Standards and Technology (NIST) (http://www.nist.gov/index.html) in the United States. To order this calibration, contact Ron Eastwood at the Roseville Standards Lab (916) 788-5506.
22 | 8491A/B Option 006/H33, 8491A/B Option 010/H33
23 | 11667A H86, 11667B H86, 11667C H86
24 | N/A
25 | Please refer to the "N7852A - Power Sensors" tab for details on Power Sensor calibration requirements for the N7852A application.
26 | H67 calibration is required for the V8486A+V281A (both need to sent in together for cal.)
27 | Only 1dB Step Atten calibration is required.
28 | H87 support extened Freq to 33GHz for N8485A option 033.
29 | When ordering periodic calibration for instruments used as lab standards in the N7800A software, we recommend using “Keysight calibration + uncertainties” for power sensors, and “Keysight calibration + uncertainties + guardbanding” for all other items. (please visit: www.Keysight.com/find/americas_cal) The N7800A software incorporates the ISO GUM meas. Uncertainty in point-to-point uncertainty calculations. The special “H-series” calibration options in this table provide lower measurement uncertainties through use of direct comparison to devices directly characterized by NPL or NIST (or another NMI). Please order option H99 to get data on a CD for easy import into the N7800A (avoids manual entry). The overall resulting N7800A measurement uncertainties then reflects these lower device uncertainties. (H99 is ONLY FOR POWER SENSORS)
30 | E9300A, E9301A, E9304A
31 | The H52 Keysight Accredited Calibration at 1 MHz is intended for an attenuators 8496G/H only. The TME calibration software employs the 8496G/H attenuator to measure @ 1 MHz. The H52 Keysight Accredited Calibration at 1 MHz is intended for an attenuators 8496G/H only. The TME calibration software employs the 8496G/H attenuator to measure @ 1 MHz.
32 | E9304A-H20 - Frequency range 9 kHz to 18 GHz, -60 to +30 dBm, APC 3.5mm connector
33 | 8485A: Required: Calibrated down to 10MHz to 26.5GHz
Note on P12: Dwayne Smallen: Footnote #33 Required: Calibrated down to 10MHz to 26.5GHz

## 11667A-H85
11667A Power Splitter Cal Option - H85 - Required Applications: N7850A, N7851A | [6 empty] | H85 - Includes electronic data file and the following specifications:
* Table depicts which and Freq Range required. |  | Power Sensor UUT - N7850A (Sensor special calibration data)
Frequency (GHz) | 11667A | N8481A, 8481A * | N8482A* | 8482A* | E9304A* | E9304A Opt H18* | Power Splitter | Source Match | Must be >0 and ≤0.5 at all points | All tests that require a power splitter
9e-06 | X |  |  |  | X | X | (Port 2 Only) | The magnitude of the equivalent source match
3e-05 | X |  |  |  | X | X |  | Source Match Phase (Degree) | Must be ≥–180° and ≤180° at all points | All tests that require a power splitter
5e-05 | X |  |  |  | X | X |  | The phase of the equivalent source match, as a degree | Must not be 0° (the default) at all points
0.0001 | X |  | X | X | X | X |  | Source Match Uncertainty | Must be non-zero | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.0003 | X |  | X | X | X | X |  | The uncertainty of the equivalent source match magnitude
0.0005 | X |  | X | X | X | X |  | Source Match Phase Uncertainty (Degree) | Must be >0° and ≤180° at all points | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.001 | X |  | X | X | X | X |  | The uncertainty of the equivalent source match phase, as a degree
0.003 | X |  | X | X | X | X
0.005 | X |  | X | X | X | X
0.01 | X | X | X | X | X | X |  | Include field that reflects which applications need this info: N7850A, N7851A
0.03 | X | X | X | X | X | X
0.05 | X | X | X | X | X | X
0.1 | X | X | X | X | X | X
0.3 | X | X | X | X | X | X
0.5 | X | X | X | X | X | X
0.8 | X |  |  |  | X | X
1 | X | X | X | X | X | X
1.2 | X |  |  |  | X | X
1.5 | X | X | X | X | X | X
2 | X | X | X | X | X | X
2.5 | X |  | X | X
3 | X | X | X | X | X | X
3.5 | X |  | X
4 | X | X | X | X | X | X
4.2 | X |  | X | X
5 | X | X | X |  | X | X
6 | X | X | X |  | X | X
7 | X | X |  |  |  | X
8 | X | X |  |  |  | X
9 | X | X |  |  |  | X
10 | X | X |  |  |  | X
11 | X | X |  |  |  | X
12 | X | X |  |  |  | X
12.4 | X | X |  |  |  | X
13 | X | X |  |  |  | X
14 | X | X |  |  |  | X
15 | X | X |  |  |  | X
16 | X | X |  |  |  | X
17 | X | X |  |  |  | X
18 | X | X |  |  |  | X

## 11667B-H85
11667B Power Splitter Cal Option - H85 - Required Applications: N7850A, N7851A |  |  |  | H85 - Includes electronic data file and the following specifications:
Frequency (GHz) | 11667B - H85 | N8485A | 8485A, 8485D | Power Splitter | Source Match | Must be >0 and ≤0.5 at all points | All tests that require a power splitter
0.01 | X | X |  | (Port 2 Only) | The magnitude of the equivalent source match
0.03 | X | X |  |  | Source Match Phase (Degree) | Must be ≥–180° and ≤180° at all points | All tests that require a power splitter
0.05 | X | X | X |  | The phase of the equivalent source match, as a degree | Must not be 0° (the default) at all points
0.1 | X | X | X |  | Source Match Uncertainty | Must be non-zero | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.3 | X | X | X |  | The uncertainty of the equivalent source match magnitude
0.5 | X | X | X |  | Source Match Phase Uncertainty (Degree) | Must be >0° and ≤180° at all points | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.8 | X | X | X |  | The uncertainty of the equivalent source match phase, as a degree
1 | X | X | X
1.2 | X | X | X
1.5 | X | X | X |  | Include field that reflects which applications need this info: N7850A
2 | X | X | X
3 | X | X | X
4 | X | X | X
5 | X | X | X
6 | X | X | X
7 | X | X | X
8 | X | X | X
9 | X | X | X
10 | X | X | X
11 | X | X | X
12 | X | X | X
12.4 | X | X | X
13 | X | X | X
14 | X | X | X
15 | X | X | X
16 | X | X | X
17 | X | X | X
18 | X | X | X
19 | X | X | X
20 | X | X | X
20.5 | X | X | X
21 | X | X | X
21.5 | X | X | X
22 | X | X | X
22.5 | X | X | X
23 | X | X | X
23.5 | X | X | X
24 | X | X | X
25 | X | X | X
26 | X | X | X
26.5 | X | X | X

## 11667C-H85
11667C Power Splitter Cal Option - H85 - Required Applications: N7850A & N7851A |  |  | H85 - Includes electronic data file and the following specifications:
Frequency (GHz) | 11667C | N8487A, 8487A | Power Splitter | Source Match | Must be >0 and ≤0.5 at all points | All tests that require a power splitter
0.05 | X | X | (Port 2 Only) | The magnitude of the equivalent source match
0.1 | X | X |  | Source Match Phase (Degree) | Must be ≥–180° and ≤180° at all points | All tests that require a power splitter
0.3 | X | X |  | The phase of the equivalent source match, as a degree | Must not be 0° (the default) at all points
0.5 | X | X |  | Source Match Uncertainty | Must be non-zero | Tests that require a power splitter are also guard-banded by measurement uncertainty
1 | X | X |  | The uncertainty of the equivalent source match magnitude
1.5 | X | X |  | Source Match Phase Uncertainty (Degree) | Must be >0° and ≤180° at all points | Tests that require a power splitter are also guard-banded by measurement uncertainty
2 | X | X |  | The uncertainty of the equivalent source match phase, as a degree
3 | X | X
4 | X | X
5 | X | X |  | Include field that reflects which applications need this info: N7850A
6 | X | X
7 | X | X
8 | X | X
9 | X | X
10 | X | X
11 | X | X
12 | X | X
12.4 | X | X
13 | X | X
14 | X | X
15 | X | X
16 | X | X
17 | X | X
18 | X | X
19 | X | X
20 | X | X
21 | X | X
22 | X | X
23 | X | X
24 | X | X
25 | X | X
26 | X | X
26.5 | X | X
27 | X | X
28 | X | X
29 | X | X
30 | X | X
31 | X | X
32 | X | X
33 | X | X
34 | X | X
34.5 | X | X
35 | X | X
36 | X | X
37 | X | X
38 | X | X
39 | X | X
40 | X | X
41 | X | X
42 | X | X
43 | X | X
44 | X | X
45 | X | X
46 | X | X
47 | X | X
48 | X | X
49 | X | X
50 | X | X

## N7852A-11667A-B-C-H86
11667X Power Splitter Cal Option - H86 - Supported Applications: N7852A (Required), N7850A & N7851A (Supported) | [9 empty] | H86 - Includes electronic data file and the following specifications:
Frequency GHz | 11667A | 11667B | 11667C | N8482A | 8482A | N8481A, 8481A | N8485A | 8485A | N8487A, 8487A | Power Splitter | Source Match | Must be >0 and ≤0.5 at all points | All tests that require a power splitter
9e-06 | X | [8 empty] | (Port 2 Only) | The magnitude of the equivalent source match
3e-05 | X | [9 empty] | Source Match Phase (Degree) | Must be ≥–180° and ≤180° at all points | All tests that require a power splitter
5e-05 | X | [9 empty] | The phase of the equivalent source match, as a degree | Must not be 0° (the default) at all points
0.0001 | X |  |  | X | X | [5 empty] | Source Match Uncertainty | Must be non-zero | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.0003 | X |  |  | X | X | [5 empty] | The uncertainty of the equivalent source match magnitude
0.0005 | X |  |  | X | X | [5 empty] | Source Match Phase Uncertainty (Degree) | Must be >0° and ≤180° at all points | Tests that require a power splitter are also guard-banded by measurement uncertainty
0.001 | X |  |  | X | X | [5 empty] | The uncertainty of the equivalent source match phase, as a degree
0.003 | X |  |  | X | X
0.005 | X |  |  | X | X
0.01 | X | X |  | X | X | X | X |  |  |  | Include field that reflects which applications need this info: N7852A
0.03 | X | X | X | X | X | X | X | X | X |  | The minimum frequency specified for 8485A, 8487A, and N8487A is 50 MHz. But the
0.05 | X | X | X | X | X | X | X | X | X |  | ETE sensor to be calibrated down to 30 MHz.
0.1 | X | X | X | X | X | X | X | X | X |  | Minimum frequency specified for N5532x sensor is 30 MHz thus needs
100 MHz step | X | X | X | X | X | X | X | X | X |  | Service center has this practice so it is ok.
4.2 | X | X | X | X | X | X | X | X | X
100 MHz step | X | X | X | X |  | X | X | X | X
6 | X | X | X | X |  | X | X | X | X
100 MHz step | X | X | X |  |  | X | X | X | X
18 | X | X | X |  |  | X | X | X | X
100 MHz step |  | X | X |  |  |  | X | X | X
26.5 |  | X | X |  |  |  | X | X | X
100 MHz step |  |  | X | [5 empty] | X
50 |  |  | X | [5 empty] | X

## 11667B-H87
11667B Power Spiltter Cal Option - H87 - Required Application: N7850A (N8485A-033) |  |  | H87 - Includes electronic data file and the following specifications:
Frequency (GHz) | 11667B - H87 | Power Splitter | Source Match | Must be >0 and ≤0.5 at all points | All tests that require a power splitter
0.01 | X | (Port 2 Only) | The magnitude of the equivalent source match
0.03 | X |  | Source Match Phase (Degree) | Must be ≥–180° and ≤180° at all points | All tests that require a power splitter
0.05 | X |  | The phase of the equivalent source match, as a degree | Must not be 0° (the default) at all points
50MHz Steps | X |  | Source Match Uncertainty | Must be non-zero | Tests that require a power splitter are also guard-banded by measurement uncertainty
2 | X |  | The uncertainty of the equivalent source match magnitude
100MHz steps | X |  | Source Match Phase Uncertainty (Degree) | Must be >0° and ≤180° at all points | Tests that require a power splitter are also guard-banded by measurement uncertainty
4.2 | X |  | The uncertainty of the equivalent source match phase, as a degree
200MHz Steps | X
12.4 | X
12.75 | X |  | Include field that reflects which applications need this info: N7850A
250MHz Steps | X
18 | X
500MHz Steps | X
26.5 | X
27 | X
1Ghz Steps | X
33 | X

## N7852A - Power Senosrs
N7852A Power Sensor Calibration Specifications:
Instrument | Critical Specifications | Recommended | Alternative
|  | Keysight Model Number | Keysight Model Num
Power Meter | Frequency Range: 100 kHz to 50 GHz | E4419B | N1912A
Absolute Accuracy: ± 0.5%
Power Reference Accuracy: ± 0.9%
Compatible with Keysight N8480 series and 8480 power sensors
Power Sensor | Frequency Range: 100 kHz to 4.2 GHz | N8482A Opt STD or CFT | 8482A
(For Option 504 only, 3 required)3
Power Sensor | Frequency Range: 10 MHz to 18 GHz | N8481A Opt STD or CFT | 8481A
(For Option 518 only, 3 required)3
Power Sensor | Frequency Range: 10 MHz to 26.5 GHz | N8485A Opt STD or CFT | 8485A
(For Option 526 only, 3 required)3
Power Sensor | Frequency Range: 30 MHz to 50 GHz | 8487A Opt STD or CFT
(For Option 550 only,3 required)3, 4
3. For each power sensor model, three units are required to be used as the test equipment below: Standard power sensor Verification power sensor Buried power sensor
4. 8487A power sensor frequency range specified in the data sheet starts from 50 MHz. To calibrate N5532A Option 550, the 8487A sensor needs to be calibrated from 30 MHz.
Device | Parameter Name and Description | Limits | A Failure Affects:
Power Sensor | Cal Factor (%) | Must be ≥10% and ≤150% at all points | All tests that require a power sensor
| The calibration factor, as a percentage | Must not be 100% (the default) at all points
| Uncertainty (%)1 | Must be >0 and ≤10%at all points | Tests that require a power sensor are also guard-banded by measurement uncertainty
The uncertainty of the calibration factor, as a percentage
| Reflection Coefficient Magnitude | Must be >0 and <0.5 at all points | Tests that require a power sensor are also guard-banded by measurement uncertainty
The magnitude of the reflection coefficient
| Reflection Coefficient Phase (Degree) | Must be ≥-180° and ≤180° at all points | All tests that require a power sensor, except for Linearity test
| The phase of the reflection coefficient, as a degree | Must not be 0° (the default) at all points
| Reflection Coefficient Uncertainty | Must be non-zero at all points | Tests that require a power sensor are also guard-banded by measurement uncertainty
The uncertainty of the reflection coefficient magnitude
| Reflection Coefficient Phase Uncertainty(+/- Degree) | Must be >0° and ≤180° at all points | Tests that require a power sensor are also guard-banded by measurement
The uncertainty of the reflection coefficient phase, as a degree
Required Calibration Data
When running a test that required the calibration data, the test will check if all the required calibration data exists. If it is not, an error message will be prompted to ask you to enter the data.
Table 1 lists the UUT factory predefined frequency points at which the calibration data is required for the external test equipment (ETEs) for each UUT model.
Table 2 lists the equipments that need to be calibrated at the specific frequency points listed in Table 1.
Please have your ETEs calibrated at those frequency points and enter the calibration data in TME before you run any test.
Table 1
Option 504 | Option 518 | N5532A/B Option 526 | U5532C Option 526 | Option 550
Frequency GHz | Frequency GHz | Frequency GHz | Frequency GHz | Frequency GHz
0.0001
0.0003
0.0005
0.001
0.003
0.005
0.01 | 0.01 |  | 0.01
0.03 | 0.03 | 0.03 | 0.03 | 0.03
0.05 | 0.05 | 0.05 | 0.05 | 0.05
0.1 | 0.1 | 0.1 | 0.1 | 0.1
100 MHz step | 100 MHz step | 100 MHz step | 100 MHz step | 100 MHz step
4.2 | 18 | 26.5 | 50 | 50
Table 2
Equipment Type | Performance Test
| VSWR | Calibration Factor
|  | N5532A | N5532A | N5532A | N5532A
|  | Option 504 | Option 518 | Option 526 | Option 550
StandardPowerSensor1 |  | *range2 | * | * | *
StandardPowerSensor2 |  | *range1
VerificationPowerSensor1 |  | *range2 | * | * | *
VerificationPowerSensor2 |  | *range1
* Refer to Table 1 for the required frequency points
range 1: 100 kHz - 50 MHz
range 2: 50 MHz - 4.2 GHz

## 8486G-H-H52
8496G and 8486H - H52 Information
Detailed Description The 8496G H52 calibration provides (Atten and VSWR at 1MHz) as an Accredited 17025 Calibration for the 8496G (10dB Steps Calibration). A data printout and an electronic copy in .csv file format on a CD-ROM, is to be provided with this option. The H52 Keysight Accredited Calibration at 1 MHz is intended for an attenuators 8496G/H only. The TME calibration software employs the 8496G/H attenuator to measure @ 1 MHz. The units require VSWR: at 1 MHz: <=1.05:1. If the units do not meet this criteria, they may be returned to the customer un-calibrated or reworked at the customer's expense, upon approval from the customer. H52 option is for a single 8496G or 8496H attenuator. The H52 calibration cannot be performed on an existing 8496G/H H50 attenuator pair.

## 8494-6-G-H - H50
TME Step Attenuator Calibration – Option H50
Table 1.
| 8494G/H Data |  |  |  |  | Specifications:
| Frequency | Step (dB) | Attenuation | Uncertainty | Reflection
| (Hz) |  | (dB) | (+/- dB) | Size | Device | Parameter Name and Description | Limits | A Failure
|  | 0 dB (Ref) | 0 | 0 |  |  |  |  | Affects:
|  | 1 dB |  | < 0.010 |  | Step Attenuator | Atten(dB) | Must be within 1 dB of the nominal attenuation. | All test runs
|  | 2 dB |  | < 0.010 |  | 8494G/H | The attenuation, in dB | Must not all be exactly equal to the nominal attenuation (the default) at all points except zero, zero step is the reference and must be zero
|  | 3 dB |  | < 0.010 |  | 8496G/H
|  | 4 dB |  | < 0.010
|  | 5 dB |  | < 0.010 |  |  | Uncert(dB) |  | Test runs guard banded by measurement uncertainty only
|  | 6 dB |  | < 0.010 |  |  | The uncertainty of the attenuation, in dB
|  | 7 dB |  | < 0.010 |  |  |  | See Table 1. For Specification
|  | 8 dB |  | < 0.010
|  | 9 dB |  | < 0.010 |  |  | ReflSize | Must be > 0 and < 0.5 at all points (TME Limit Checker) | Test runs guard banded by measurement uncertainty only
|  | 10 dB |  | < 0.010 |  |  | The reflection coefficient magnitude
|  | 11 dB |  | < 0.010
| 8496G/H Data | [5 empty] | VSWR: | at 50 MHz: £1.05:1
| Frequency | Step (dB) | Attenuation | Uncertainty | Reflection
| (Hz) |  | (dB) | (+/- dB) | Size
|  | 0 dB (Ref) | 0 | 0
|  | 10 dB |  | < 0.010 |  | Applicable Models
|  | 20 dB |  | < 0.015 |  | 8494/8496G, Type N Input
|  | 30 dB |  | < 0.020 |  | 8494/8496H, Type N Input
|  | 40 dB |  | < 0.025
|  | 50 dB |  | < 0.030 |  | Description:
|  | 60 dB |  | < 0.035 |  | The TME calibration software employs the 8494G/H and 9496G/H attenuators to measure Scale Fidelity, Attenuator Accuracy, and Absolute Amplitude Accuracy @ 50 MHz.
|  | 70 dB |  | < 0.040 |  | This document describes the calibration requirements for the attenuator set. The attenuators will be permanently connected using the 11716A Interconnect kit as shown in Figure 1.
|  | 80 dB |  | < 0.046
|  | 90 dB |  | < 0.055
|  | 100 dB |  | < 0.065
|  | 110 dB |  | < 0.075
Data Provided:
A data printout and an electronic copy in .csv file format on a
USB drive or CD-ROM (Customer selected), is to be provided with this option.
8494G/H Settings
| Attenuation | S1 | S2 | S3 | S4
| Setting |  |  |  |  | Figure 1: TME Step Attenuator
| 0 dB (Ref setting) | Off | Off | Off | Off
| 1 dB | On | Off | Off | Off | The 8494/8496 series attenuators are made up of four attenuator sections. The attenuator sections are arranged in a 1/2/4/4 dB, or 10/20/40/40 dB format. The 8494G/H Step Attenuator can stepped from 0 dB to 11 dB in 1 dB steps. The 8496G/H Step Attenuator can be stepped from 0 dB to 110 dB in 10 dB steps.
| 2 dB | Off | On | Off | Off | For both the 8494, or 8496 attenuators the selection of 4 dB, or 40 dB of attenuation can be done with either Section 3, or Section 4. The Production Line step attenuators use Section 4, and the TME software uses Section 3. The TME attenuators have been characterized in the Roseville Standards Lab and this process has worked well until 2007.
| 3 dB | On | On | Off | Off | Starting in 2007 Keysight began to provide ISO17025, and Z540 Cals on new products from the factory . The quickest way to provide these cals was to setup a TME test station at the end of the production line. Having a TME test station within the factory came the need to characterize TME Step Attenuators in the Factory Metrology Labs. This document will describe the Calibration requirements for the TME Step Attenuators.
| 4 dB | Off | Off | On1 | Off | The calibration should be performed at 50 MHz, however other frequencies could be necessary on future products.
| 5 dB | On | Off | On1 | Off
| 6 dB | Off | On | On1 | Off
| 7 dB | On | On | On1 | Off
| 8 dB | Off | Off | On | On
| 9 dB | On | Off | On | On
| 10 dB | Off | On | On | On
| 11 dB | On | On | On | On
8496G/H Settings
| Attenuation | S5 | S6 | S7 | S8
Setting
| 0 dB (Ref setting) | Off | Off | Off | Off
| 10 dB | On | Off | Off | Off
| 20 dB | Off | On | Off | Off
| 30 dB | On | On | Off | Off
| 40 dB | Off | Off | On1 | Off
| 50 dB | On | Off | On1 | Off
| 60 dB | Off | On | On1 | Off
| 70 dB | On | On | On1 | Off
| 80 dB | Off | Off | On | On
| 90 dB | On | Off | On | On
| 100 dB | Off | On | On | On
| 110 dB | On | On | On | On
1) TME requires that Section 3 be characterized.

## 8491A-B-006-010-H33
8491A-B-006-010 - H33 Option Specifications:
| Attenuators | Critical Specs | Recommended | Alternate
| 6 dB Fixed Attenuator | Frequency: 50 MHz | 8491A/B Option 006/H33 | 8491A/B/C Option 006
|  | VSWR: < 1.05:1 |  | See Footnote a
Loss: 6 dB (nominal)
| 10 dB fixed Attenuator | Frequency: 50 MHz | 8491A/B Option 010/H33 | 8491A/B/C Option 010
|  | VSWR: < 1.05:1 |  | See Footnote: a
Loss: 10 dB (nominal)
Footnote a) The standard 8491A VSWR is specified at <1.2:1 , DC to 8 GHz. Most 8491A/B/C attenuators will meet a VSWR of <1.05:1 at 50 MHz. The VSWR should be calibrated on an periodic basis.